Technical Specifications :

Analyzer wafers for XES experiment at SP12U

National Synchrotron Radiation Research Center

Hsinchu, Taiwan R.O.C.

(NOV, 2025)

CONTENTS

1	SCOPE OF THE ORDER

2	CRYSTALS

3	INSPECTION AND ACCEPTANCE CRITERIA

4	DELIVERY

- SCOPE OF THE ORDER

The scope of the present order is to manufacture (i) circular Si crystal wafers for the x-ray emission spectrometers at 12U. The wafers have a surface normal to the [531] or [553] direction. The diameter of the wafer is 100-mm while the thickness 0.5-mm for all the wafers. The double side needs to be mirror polished.

- Crystals

The 6 wafers should have the reflection surface of [531] and 6 wafers of [553]. The wafers are lapped and etched for all surfaces to eliminate any visible scratches, point defects and then strain free. The double sides should be mirror polished. The orientations of the crystals are shown in Fig.1. The detail specifications of crystals are shown below. The orientation flat or notch is accepted.

| a | Material | Floating zone dislocation free Si crystal Resistivity > 1000 Ω-cm |  |  |
| --- | --- | --- | --- | --- |
| b | Diffraction Plane | 531 | 553 |  |
| c | Quantity | 6 wafers | 6 wafers |  |
| d | Dimension (except for flat/notch) | 100 (< ± 0.5) mmφ 0.5 (< ± 0.05) mm t | 100 (< ± 0.5) mmφ 0.5 (< ± 0.05) mm t |  |
| e | Angular accuracies | < 0.5 deg | < 0.5 deg |  |
| f | Polishing | Mirror polishing, Double sides |  |  |

Table 1. Requirements for the wafers

- INSPECTION AND ACCEPTANCE CRITERIA

  - Visible inspection

Visible damages, such as clacks or dips of > 0.1 mm, are not accepted for the reflection surfaces, while those of > 1 mm are not for other surfaces, except for marks to identify the crystal orientation. The refection surfaces have mirror-like surfaces.

  - Material report

The supplier should provide NSRRC with the material report for the Si ingot. The ingots have to be floating-zone Si crystal with a resistivity of > 1000 Ω-cm [see, Table-I (a)].

  - Factory acceptance test (FAT) report

The supplier should provide NSRRC with the factory acceptance test (FAT) report for the wafers [For the check items, see Table-I(b-f)].

- DELIVERY

The crystals should be delivered in 270 calendar days by the date of awarding contract. The supplier is responsible for the delivery of all the components, including the payment of the import/consumption tax. The address for the delivery is BL12XU at SPring-8, 1-1-1 Kouto, Sayo-cho, Sayo-gun, Hyogo 679-5198, Japan (DDP, Hyogo, JP, INCOTERM2020)

Fig .1 Required dimensions and orientations of the circular wafers