Specifications of XBPM vacuum chamber

Procurement number:

Designation of procurement : Specifications of XBPM vacuum chamber

The National Synchrotron Radiation Research Center(hereinafter referred to as "NSRRC") shall be pleased to receive full details of your quotation in accordance with the following terms and conditions：

Overviews

The document is for the bid of the XBPM vacuum chamber, including a high vacuum (HV) compatible flange and motions at the TPS 14A beamline at Taiwan Photon Source in Hsinchu, Taiwan.

Scope of the contract

The vendor should provide:

- XBPM vacuum chamber, including a high vacuum (HV) compatible flange and horizontal and vertical motions

- Factory test report.

- User manual.

- Technical description

The single-crystal diamond beam position monitors (XBPM®), developed by CIVIDEC (Austria), offer high-precision X-ray and neutron diagnostics. Their sCVD XBPM® detector achieves position resolution of 10⁻³ of the X-ray beam size. The thin diamond layer and titanium electrodes reduce wavefront distortion compared to polycrystalline detectors.

The EPICS-compatible XANDY system includes an integrated power supply and four low-noise, high-gain transimpedance amplifiers (1 pA–1 mA input range) for XBPM® signals (A–D). It enables autonomous beam monitoring and feedback control for accelerators or optical components, with protection against over-voltage from asymmetric beam loads.

The diamond sensor is mounted inside a high vacuum-compatible flange with sub-micrometre-resolution horizontal and vertical motions. Stepper motors are externally fixed and operate via magnetically coupled feedthroughs, ensuring vacuum isolation and stable sensor positioning.

The main parts of an XBPM Monitor are:

1. 20 µm thick single crystal Diamond XBPM® Detector with titanium electrodes

2. EPICS compatible X-Ray Beam Position Monitoring System XANDY

3. An HV-compatible flange-mounted casing

4. Horizontal and vertical motions inside the casing

For a detailed description of items 1 and 2, especially their performance, see the dedicated information by CIVIDEC. This offer concerns items 3. and 4

- 1 Description of System Components and Coordinate System Labels:

The XBPM vacuum chamber must include the following main components: HV-compatible flange-mounted casing and motions with two sets of orthogonal rails.

The CIVIDEC B9 XBPM diamond sensor will be placed inside a HV-compatible flange-mounted casing (DN160 CF flange), which is a system capable of horizontal and vertical (X and Z) motions with sub-micrometre resolution. The interface with the beamline will be a KF-40 flange (all other flanges will be CF type).

The mechanical drawing and connector layout of the CIVIDEC B9 XBPM diamond sensor are shown in Figure 1 below:
(a) The mechanical dimensions of the sensor, including mounting holes, measurement aperture, parking position, and section views with detailed tolerances.
(b) The connector layout, showing the SMA connectors (SENSOR 1–4) and the BNC connector (BIAS), along with their respective feedthroughs and cabling directions. The beam direction is indicated as pointing into the page.

(a)

(b)

Figure 1: (a) the mechanical drawing and (b) connector layout of the CIVIDEC B9 XBPM diamond sensor

1.2 Coordinate System Description (Figure 2):

  Y-axis: Parallel to the X-ray beam direction.

  X-axis: Perpendicular to the X-ray beam direction and parallel to the ground.

  Z-axis: Perpendicular to the ground.

Figure 2: Coordinate System Description

- 3 The Specification of XBPM Vacuum Chamber:

    - Specification of an HV-compatible flange-mounted casing

The conceptual drawings of the system are shown in Figures 1(a) and 1(b), and its specifications are listed below:

- Items 1–10 are described in Table 1, including material and technical descriptions.

- CF160 flanges are used as the main vacuum interface, with adapters (item 10) converting from CF160 to KF40 flanges, aligned along the X-ray beam direction as shown in Figure 1(b).

- The system will be sealed with an O-ring and is capable of reaching a base pressure at or below 10−5 bar.

- The feedthroughs of items 1–4 are DN16CF SMA connectors (model KVFC2031), with female SMA connectors on both the inside and outside of the casing, requiring SMA male plug cables for connection.

- The feedthrough of item 5 is a DN16CF BNC connector, with female BNC connectors on both the inside and outside of the casing, requiring BNC male plug cables for connection.

(a)

(b)

Figure 3: Conceptual drawings of the HV-compatible flange-mounted casing: (a) Top view along the X-ray beam direction, showing the internal components and feedthrough positions; (b) Side view, showing the beam direction and the CF160-to-KF40 adapter flange (item 10).

| item | Description | materials |
| --- | --- | --- |
| 1-4 | DN16CF SMA FEEDTHROUGH- Connector type (SMA) | 304L SS |
| 5 | FEEDTHRU BNC CF16- Connector type (BNC) | 304L SS |
| 6 | Horizontal Motors:2-phase | 304L SS |
| 7 | Vertical Motors: 2-phase | 304L SS |
| 8 | DN160CF (8.00" OD) | 304L SS |
| 9 | Base plate |  |
| 10 | CF 160 to KF 40 | 304L SS |

Table 1: The specification of an HV-compatible flange-mounted casing

  - 3 Specification of motions

The in-vacuum motions are guided by two sets of orthogonal rails. The horizontal motion will be equipped with two runner blocks to minimize stick-slip. Furthermore, the motions will be equipped with normally closed limit switches, and the housing itself will function as a hard stop. UHV-compatible Renishaw absolute encoders communicating via BiSS-C protocol will be used for all motions. The stepper motors are fixed to the stationary UHV casing, and the motions are established via magnetically coupled feedthroughs. This design ensures that there are no vacuum forces on the system, which results in stable positioning of the diamond sensor even when power to the motors is lost.

- The stepper motors (item 6 and 7) are fixed to the HV casing, and the motions are established via magnetically coupled feedthroughs.

| Type of motion | 2x linear, vertical and horizontal (Z and X) |
| --- | --- |
| Motor & gearbox type | McLennan, 17HS-240E, 100:1 |
| Guidance in a vacuum | Fine thread spindle drive and cross roller bearing |
| Max current | 2.3A/phase |
| Resistance | 0.72Ω/phase |
| Motor step angle | 200 steps/rev, 400 steps/rev (half steps) |
| Resolution | 0.2µm |
| Repeatability | 0.5µm |
| Range | 10 mm vertically and horizontally (Z and X) |

Table 2: Motion Specification

| Summary of motions and ranges Motion; Range; Resolution; Repeatability. X translation; >=10mm; <=0.2µm; <=0.5µm. Z translation; >=10mm; <=0.2µm; <=0.5µm. |  |  |  |
| --- | --- | --- | --- |

Inspection & Acceptance Criterion:

- The drawings must be provided within 60 calendar days of the contract award. They shall meet the requirements outlined in Section 1.3 of the specification

- Factory Acceptance Report, including (1) the supplier shall provide material certification documents for materials (items 1-10), and (2) the test of motions (items 6 and 7): translation range, resolution, and repeatability (3) vacuum test.

Delivery time:  182 calendar days

Payment schedule: NSRRC shall open an irrevocable Letter of Credit.