Technical Specifications :

Closed cycle cryostat for X-ray diffraction

National Synchrotron Radiation Research Center

Hsinchu, Taiwan R.O.C.

(May 20, 2026)

CONTENTS

1	SCOPE OF CONTRACT

2	GENERAL SYSTEM REQUIREMENTS

3	INSPECTION AND ACCEPTANCE CRITERIA

4	DELIVERY

5	RESPONSIBILITIES OF THE SUPPLIER

- SCOPE OF CONTRACT

The objective of this specification is the replacement and upgrade of the current x-ray diffraction experiment cryostat system. The new system is expected to provide enhanced thermal stability, reduced vibration, and improved X-ray transparency to support advanced structural analysis, including Pair Distribution Function (PDF) analysis. The system shall be a turnkey, "dry" (cryogen-free) solution comprising the cryostat, a high-performance temperature controller, and an integrated vacuum pumping station.

- GENERAL SYSTEM REQUIREMENTS

- General

  - Cooling Engine: Closed-cycle refrigerator.

  - Temperature Range: Continuous operation from 50 K to 700 K without hardware modifications.

  - Vibration Isolation: The system must incorporate a vibration-damping interface (e.g., bellows-isolated vacuum shroud and flexible thermal links) to ensure the sample remains stable within the X-ray beam.

- Control and Sensors

  - Control Loops: At least two independent PID control loops.

  - Remote Interface: Integrated LAN (Ethernet) port for remote monitoring and control.

  - Communication Protocols: Support for standard protocols (e.g., TCP/IP, HTTP, or VXI-11) to ensure compatibility with LabVIEW, Python, and EPICS-based beamline control systems.

  - Sensors: Dual-sensor configuration optimized for the full range (e.g., Silicon Diode for < 300 K and PT100/Type-E Thermocouple for > 300K

- Sample Environment

  - X-ray Windows: Low-absorption windows.

  - Angular Access: Radial or conical scattering access of at least 120 degrees.

  - Internal Geometry: Optimized sample holders.

- Vacuum Pump

  - Pump system: Turbomolecular pump with oil-free dry scroll pump

  - Ultimate Pressure: < 1x10-4 Pa (at room temperature).

  - Gauging: Full-range vacuum gauge integrated into the system for real-time pressure monitoring.

  - Integrated Vacuum System:　 The pumping system shall be mounted on a dedicated, independent stand equipped with heavy-duty casters for mobility.

- INSPECTION AND ACCEPTANCE CRITERIA

  - Factory acceptance test reports

The supplier should provide NSRRC with the factory acceptance test reports. The performances specified in (2.2、2.5、2.9、2.11、2.12、2.13) must be ensured.

  - On-site function test report

The supplier is responsible for the on-site function test. NSRRC will supervise the test. The lowest temperature ≦ 50 K and highest temperature ≧700K must be ensured.

3.3 For specifications not covered by the above tests, the supplier shall provide product catalogs (or technical documents/specifications) as supporting documentation to demonstrate compliance with the required specifications.

- DELIVERY

The supplier is responsible for the delivery of all the components, including the payment of the importing/consuming tax. The address for the delivery is BL12B2 at SPring-8, 1-1-1 Kouto, Sayo-cho, Sayo-gun, Hyogo 679-5198, Japan (DDP, Hyogo, JP, INCOTERM2020)

- RESPONSIBILITY OF SUPPLIER

  - The design for the cryostat and other environments in the sample space including pumping station. Several discussions are required before the details finalized with NSRRC. The supplier should provide NSRRC with the drawings within 90 calendar days after awarding the contract. The supplier should deliver the products within 300 calendar days after the drawing confirmation and noticed by NSRRC. The installation and function test should be completed within 30 calendar days after the notification from NSRRC.

  - The manufacturing and the transportation of the required products to the Taiwan Beamline Office at SPring-8 site.

  - Factory acceptance test (FAT) reports before the assembly

  - On-site function test report after the assembly.