SRPO-1150147

高速遮光器 規格書

Specification of fast photon shutter

- 概述

  - 此購案欲採購一遮光器，遮蔽光源為同步輻射光源。詳細需求規格列於下方。

- 需求規格

  - 真空相容
可操作在氣壓 1E-3 torr 至 760 torr 間。

  - 使用壽命
大於五千萬次循環

  - 開關延遲
收到訊號至動作開始的時間 < 0.2 ms

  - 遮光材質
鎢或鎢合金，遮光材料於光軸方向之物理厚度 > 2.5 mm

  - 狀態偵測
須具備遮光器開關之狀態偵測能力

  - 開關控制
可經由外部訊號直接控制

  - 遠端控制
可經由Ethernet/RS232/RS422介面進行程式遙控

  - 孔徑
> 2.5 mm

  - 開關切換時間
遮光片從全開到全關（或反之）的動作過程時間 < 10 ms

  - 外觀尺寸(不含線材)
< 80 X 80 X 80 mm3

  - 主體重量(不含線材)
< 300 g

  - 內容物
此購案為整套系統採購，須包含但不限於遮光器，驅動器，控制器主機及通訊模組，以及運轉所需所有配件。

- 交貨期限

  - 決標日起120日曆天

- 交貨地點

  - 國家同步輻射研究中心

- 保固

  - 完成履約經驗收合格日起，由廠商保固12個月

- 驗收方式

  - 如為型錄產品，採其公開規格值。未公開之規格且列於需求規格內者，需提供出廠證明。

  - 本案於驗收時，將由同步輻射研究中心指派之人員，檢視其尺寸與重量是否符合規格，並將設備與所有配件組合後，使用遠端遙控介面（如 Ethernet/ RS232 / RS422）對遮光器系統進行操作測試，以驗證系統是否能正常執行開關動作與功能。

- Overview

  - This procurement is for a high-speed shutter system designed to block synchrotron radiation light sources. The system must meet the detailed technical requirements listed below.

- Technical Requirements

  - Vacuum Compatibility:
Capable of operating within a pressure range of 1E-3 Torr to 760 Torr.

  - Service Life:
More than 50 million cycles.

  - Opening/Closing Delay:
The time from receiving a trigger signal to the start of mechanical movement must be < 0.2 ms.

  - Shutter Material & Thickness:
Made of Tungsten or Tungsten alloy, with a physical thickness of > 2.5 mm along the beam axis.

  - Status Feedback:
The system must provide real-time monitoring of the shutter’s open/closed status.

  - Trigger Control:
Direct control via external hardware signals must be supported.

  - Remote Control:
Remote programming and control must be available via Ethernet, RS232, or RS422 interfaces.

  - Aperture:
Minimum clear aperture of > 2.5 mm.

  - Switching Time:
The duration of the movement process from fully open to fully closed (or vice versa) must be < 10 ms.

  - Dimensions  (cables are not included):
Overall envelope size must be < 80 X 80 X 80 mm³.

  - Main Body Weight (cables are not included):
Total weight of the shutter head must be < 300 g

- Scope of Supply

  - This is a turnkey system procurement. The vendor must provide all components necessary for full operation, including but not limited to: shutter head, driver and controller unit, communication modules, and all required accessories.

- Delivery and Terms

  - Delivery Lead Time:
Within 120 calendar days from the day of the award of the contract.

  - Delivery Location:
National Synchrotron Radiation Research Center (NSRRC), Taiwan.

  - Warranty: 12 months starting from the date of successful final acceptance.

- Acceptance Criteria

  - For off-the-shelf catalog products, the manufacturer’s published specifications will be used for verification. For requirements not explicitly listed in public catalogs, the vendor must provide an official factory test report or certification of compliance.

  - During final acceptance, personnel designated by the National Synchrotron Radiation Research Center (NSRRC) will inspect the dimensions and weight of the equipment to ensure compliance with the specifications. After assembling the device with all required accessories, the personnel will operate the shutter system via remote control interfaces (e.g., Ethernet, RS232, or RS422) to verify that the opening, closing, and overall system functions operate normally.