購案編號：SRPO-1150327

高真空用聚焦凹面鏡

Specifications of B4C coated concave mirror

May 2026

Contents

1. INTRODUCTION..................................................................................................... 4

1.1 Purpose.....................................................................................................................4

1.2 Scope of specification ..............................................................................................4

2. REQUIREMENTS.....................................................................................................4

2.1 General....................................................................................................................4

2.2 Quantity...................................................................................................................4

2.3 Operating conditions ..............................................................................................4

2.4 Geometrical parameters ........................................................................................5

2.5 Tolerance..................................................................................................................5

2.6 Materials .................................................................................................................5

3. QUALITY ASSURANCE……………………….....................................................5

3.1 Responsibility for inspection……………...............................................................5

3.2 Grinding and polishing……………..................................................................5

3.3 Surface roughness……………..................................................................5

3.4 Factory acceptance test (FAT)……………...............................................................5

3.5 Acceptance criteria……………...............................................................6

4. PREPARATION FOR DELIVERY…………….....................................................6

4.1 Cleanliness……………..................................................................6

4.2 Packaging……………..................................................................6

4.3 Packing……………..................................................................6

1. INTRODUCTION

1.1 Purpose

This document contains the specifications for the B4C coated concave mirror at the National Synchrotron Radiation Research Center (NSRRC).

1.2 Scope of specification

This specification establishes the requirements for the fabrication, inspection, testing, and delivery of the EUV concave mirror.

2. REQUIREMENTS

2.1 General

These mirrors would be in the vacuum environment for the TPS 30A beamline at NSRRC.

2.2 Quantity

Two (2)

2.3 Operating conditions

Vacuum pressure:	< 1 x 10-6 torr

Working spectral range: 30 ~ 70 eV

2.4 Geometrical parameters

Dimensions: 101.6 mm dia. ±0.2 mm

Thickness: 25.4 mm ±0.1 mm

Surface Shape: concave

Radius of curvature: 20 m ±1%

2.5 Surface error tolerances

Substrate surface roughness: 0.5 nm RMS or less

2.6 Materials

Substrate: Fused silica

Coating: B4C (Boron Carbide)

.

3. QUALITY ASSURANCE

3.1 Responsibility for inspection

The supplier is responsible for the inspections of all performances as specified. The supplier may utilize his own facilities or any commercial laboratory acceptable to NSRRC. A suggested test for figure verification is long-trace, non-contact surface profilometry.

3.2 Grinding and polishing

The suppliers should provide a detailed description of the entire fabrication process and NSRRC reserves the right to approve it.

3.3 Surface roughness

The supplier should provide evidence that the final polished surface meets the surface finish requirements in paragraph 2.5.

3.4 Factory acceptance test (FAT) should include the following items:

- The dimension of clear aperture

- The surface roughness

- The surface quality

- The substrate surface figure error

- The substrate surface roughness

- Coating thickness

- Target energy range

3.5 Acceptance criteria

-The Seller shall submit factory acceptance test (FAT) as specified in section 3.4.

-NSRRC Functional Test

NSRRC will measure the radius of curvature and substrate surface figure error. The error should be within ± 3 %

4. PREPARATION FOR DELIVERY

4.1 Cleanliness

The finished pieces must be thoroughly clean and free from any organic low-vapor pressure materials or residual polished compound or cutting fluid.

4.2 Packaging

Each individual mirror must be packaged in its own protective container. The container must be capable of being sealed tightly to prevent dust and moisture from entering. The container must prevent contact or any object with any part of the clean aperture. Design of the packing container is subject to NSRRC approved.

4.3 Packing

Packing for shipment must ensure that each mirror is insulated from severe shock and rough handling. Packing marking shall indicate the fragile nature of the contents.

高真空用聚焦凹面鏡 購案規格 (Specification of HV concave mirrors)

產品名稱: 高真空用聚焦凹面鏡 (Product Name: HV concave mirror)

數量: 一式 (2片) (Quantities: 2 pcs)

| Parameter | Value | Note |
| --- | --- | --- |
| 基板材料 (Substrate material) | 熔融石英 (Fused silica) |  |
| 尺寸 (Dimensions) | 直徑101.6 mm | ±0.2 mm |
| 有效口徑 (Clear aperture) | 直徑91.4mm | ±0.2 mm |
| 厚度 (Thickness) | 25.4 mm | ±0.1 mm |
| 曲率半徑 (Radius of curvature) | 20 m 凹面 (concave) | ±1% |
| 基板表面形貌誤差 (Substrate surface figure error) | λ/10 @633 nm |  |
| 基板表面粗糙度 (Substrate surface roughness) | ≦0.5 nm RMS or less |  |
| 鍍膜材料 (Coating material) | 碳化硼 B4C (Boron Carbide) |  |
| 鍍膜厚度 (Coating thickness) | 50 nm | ±10 nm |
| 目標能量範圍 (Target energy range) | 30 eV – 70 eV |  |