National Synchrotron Radiation Research Center

Specification of

High heat load slits for SPring-8 12XU

(SRPO-1140457)

- Introduction

- Purchasing items

- Specification

- Test Standard-Acceptance Criteria

- Delivery

- Introduction

This document is the specification of high heat load slits for SPring-8 12XU. The symbols X and Y mentioned in this document are two orthogonal directions in the horizontal plane, and symbol Z is the vertical direction. The Y-axis is aligned with the direction of the X-ray.

- Purchasing items

- SPring-8 12XU High heat load slits, which includes:

  - Two horizontal fully overlapping blades module

  - Two vertical fully overlapping blades module

  - Thermal simulation of blades

  - Water cooling system (vacuum side, including feedthrough)

  - Vacuum chamber with CF ports

  - Chamber support with manual adjustment mechanism.

  - Turbo pump and its support.

- Specification

- Space limitation: Considering the limited space of 12XU beamline, the dimension of the high heat load slits shall NOT exceed the boundary shown in the following sketch.

Figure 1.  Space limitation of the high heat load slits

- Vacuum condition:

  - The vacuum level of the chamber shall be better than 5 x 10-6 Pa.

  - The He leakage rate shall be less than 5x10-10 mbar∙L/s.

- Flanges and ports:

  - The flanges at upstream and downstream sides are CF2.75” and CF6”, respectively.

  - The Max. distance between the upstream and downstream flange is 920 mm (face to face).

  - Extra six CF2.75” ports for gauges and electronic feedthrough shall be included.

  - One Pumping port shall be included for STP-301 DN100 or STP-iX457 DN160 turbo pump.

- Crane inside hutch: The Max. lift height and load capacity of the crane inside hutch is 2500 mm and 1 ton, respectively.

- Beam Power: The max. power and power density on the blade is 100 W and 400 W/mm2, respectively.

- Motors and Connectors

  - Motors

Five-phase stepping motors that are compatible with the existing motor drivers, MELEC H750v1/GDB-5F40 (Type Ⅱ SPDC) or H584A-01/GD-5610v1 (Type Ⅲ A). (Widely available. NSRRC can provide the vendors if not available)

  - Connectors

BURNDY 8 pin connectors must be used. (Widely available. NSRRC can provide the vendors if not available)

  - Limit switches

One limit switch for the positive direction and one for the

negative direction.

A,B,C,D,H: Power (5 phase), E: LS-COMMON, F: CW-LS, G: CCW-LS

- Travel Range of blades: All blades shall be able to move ±5 mm.

- Encoder resolution and min. step of blades: Encoder resolution for all blades shall be better than 0.1 um. The min. step size of blades shall be smaller than 0.2 um.

- Assembly tolerance of blades: The misalignment of blades shall meet the following requirement:

  - The parallelism between the two blade edges must be better than 5 mrad.

  - The perpendicularity of all jaws to the beam axis must be better than 1 mrad.

- Manual adjustment of chamber:

  - Adjustable range in X, Y, and Z direction shall be larger than ±15 mm. Max. screw pitch of adjustment unit shall be smaller than 1 mm.

  - Adjustable range for pitch, roll and yaw shall be larger than ±10 mrad. Max. screw pitch of adjustment unit shall be smaller than 1.5 mm.

- Thermal couples:

At least one thermocouple shall be installed on each blade module to monitor the temperature at critical positions during operation. The cables for thermal couple (vacuum side) shall be included.

- Water Cooling system:

Cooling pipes (on the vacuum side) and the corresponding feedthroughs shall be included.

Cooling water can be supplied either from a chiller at a pressure of 1 kg/cm² or from the SPring-8 cooling water line at 6 kg/cm²

- Drawings and approval:

  - The vender shall provide the drawing within 150 calendar days from the date of awarding contract.

  - NSRRC shall review the drawing in 60 calendar days form the date of receiving drawing.

- Blades:

  - The blades shall be supplied by NSRRC after the drawing has been approved.

- Test Standard-Acceptance Criteria

- The supplier shall test the performance of high heat load slits before shipment, and issue an approved test report. The following items shall be included in the FAT:

  - The travel range of each blade.

  - The min. step of each blade.

  - The manual adjustable range of chamber.

  - Thermal simulation results of blades

- The supplier shall issue certificate of manufacture document.

- Once the high heat load slits is delivered, NSRRC will check its function at SPring-8 12XU in 60 calendar days.

- Training: The training shall be completed within 60 calendar days from the day following the date of delivery.

- Delivery

The SPring-8 12XU High heat load slits shall be delivered to SPring-8 12XU.

- Incoterms/delivery term: DDP SPring-8 12XU unloaded

- Payment term: Letters of Credit (LC), 30 days after inspection acceptance.

- Deadline for shipment: The delivery shall be completed within 150 calendar days from the date of receiving the blades provided by NSRRC.

- Packages: The stages shall be delivered with appropriate package.

- The SPring-8 12XU High heat load slits shall be delivered to SPring-8 Taiwan Beamline Office.

Consignee : Japan Synchrotron Radiation Research Institute (JASRI)

address: 1-1-1, Kouto, Sayo-cho, Sayo-gun, Hyogo, 679-5198, JAPAN

TEL: +81-(0)791-58-0961

Notify Party :National Synchrotron Radiation Research center (NSRRC)

Taiwan Beamline Office at SPring-8 (Dr. Yu-Cheng, Shao)

address: 1-1-1, Kouto, Sayo-cho, Sayo-gun, Hyogo, 679-5148, JAPAN

TEL: +81-(0)791-58-1867)