Specification of SRPO-1140388

1. Scope of Supply

The scope of this procurement is to manufacture and deliver high-purity Si and Ge crystal wafers for use in high-resolution x-ray emission spectrometry (XES) at beamline 47A . The wafers must meet stringent quality standards in orientation accuracy, surface polish, mosaicity, and material purity.

2. Crystal Wafer Specifications

| Parameter | Specification |
| --- | --- |
| Materials | - Floating-zone (FZ), dislocation-free Si single crystal - High-purity Ge single crystal (dislocation-free) |
| Diffraction Planes | - Si: (100), (110), (111), (553), (620) |
| Quantity | 10 wafers per orientation, total 50 wafers |
| Wafer Dimensions | Diameter: 100 mm ± 0.5 mm Thickness: 0.5 mm ± 0.05 mm |
| Mosaicity | < 5 arcsec (FWHM, rocking curve) |
| Angular Accuracy | - Si(100), Si(110), Si(111): < 0.02° - Si(553), Si(620): < 0.04° |
| Surface Quality | Double-side mirror polished, lapped and etched |
| Resistivity | - Si: > 1000 Ω·m (FZ grown) |

3. Inspection and Acceptance Criteria

3.1 Visible Inspection

Visible damages, such as clacks or dips of > 0.1 mm, are not accepted for the reflection surfaces, while those of > 1 mm are not for other surfaces, except for marks to identify the crystal orientation. The refection surfaces have mirror-like surfaces.

3.2 Factory Acceptance Test (FAT)
- FAT report must include mosaicity, orientation accuracy, visual inspection results.

3.3 Material Certificate
- Si: FZ-grown, resistivity > 1000 Ω·m.

4. Delivery

Si wafers (including all orientations: Si(100), Si(110), Si(111), Si(553),Si(620)) shall be delivered within 320 calendar days by the date of awarding contract.

Fig .1 Required dimensions and orientations of the circular wafers

SRPO-1140388規格書

1. 供應範圍

本採購案旨在製造並交付高純度矽與鍺晶體晶圓，供47A光束線之高解析度X光發射光譜儀(XES)使用。晶圓需符合晶面角度、表面拋光、鑲嵌性與材料純度等高規格品質要求。

2. 晶體晶圓技術規格

| 項目 | 規格 |
| --- | --- |
| 材料 | Floating-zone (FZ), dislocation-free Si single crystal High-purity Ge single crystal (dislocation-free) |
| 晶面 | - 矽：(100)、(110)、(111)、(553)、(620) |
| 數量 | 每種晶面10片，共50片 |
| 晶圓尺寸 | 直徑100 ± 0.5 mm，厚度0.5 ± 0.05 mm |
| 鑲嵌性 | 小於5 arcsec |
| 晶面準確度 | - Si(100)、Si(110)、Si(111)、< 0.02° Si(553)、Si(620)：< 0.04° |
| 表面品質 | 雙面鏡面拋光，拋光與腐蝕處理以去除缺陷與殘餘應力 |
| 電阻率 | - 矽：> 1000 Ω·m |

3. 檢驗與驗收標準

3.1 目視檢查

反射面若出現大於 0.1 mm 的裂痕或凹痕等可見損傷，則不予接受；其他表面若有大於 1 mm 的損傷，同樣不被接受，唯有標示晶體取向的記號可作為例外。
反射面須具有鏡面般的光滑外觀。

3.2 工廠驗收測試 (FAT)
- 需提供工廠測試報告，內容包括：鑲嵌性、角度偏差、目視檢查等項目。

3.3 材料報告
- 矽：需為FZ製程且電阻率 > 1000 Ω·m。

4. 交貨日期

矽晶圓（包含所有晶面方向：Si(100)、Si(110)、Si(111)、Si(553)、Si(620)）須於契約決標日起 320 個日曆天內完成交付。

圖 1：所需圓形晶圓之尺寸與晶面方向。