# Specifications of a deflection mirror chamber at TPS 14A beamline

National Synchrotron Radiation Research Center

101 Hsin-Ann Rd., Hsinchu 30076, Taiwan

2025年10月23日

The National Synchrotron Radiation Research Center (hereinafter referred to as "NSRRC")

shall be pleased to receive full details of your quotation in accordance with the following terms and conditions:

# Overviews

The document is for the bid of a delection mirror chamber at TPS 14A beamline at Taiwan Photon Source in Hsinchu Taiwan.

# Scope of the contract

The vendor should provide:

- One mirror chamber including mirror holders and mirror adjustments.

- Electronic cabling.

- Factory test report.

- User manual.

- One screen.

- On-site installation

# Technical description

TPS 14A beamline will be dedicated to X-ray small angle scattering experiments. The source is a 1.19T bending magnet. A double multilayer monochromator(DMM) is placed at 23m from the source and used to select the working energy. After the DMM, a toroid focusing mirror is placed at 26m. It images the source to 50m which is at the surface of the detector. The sample position is 45m from the source.  Between the TFM and the sample, a plane mirror(PM) is placed at 27.2m. The plane mirror bends the output beam of the TFM and makes the monochromatic beam parallel to the ground. Considering the rejection of the high harmonics, the PM has two reflection stripes. One is Rh and the other is bare-silicon. Finally, a deflection mirror(DM) is used to bend the beam downward for some specific experiments. The center of the DM is 44m from the source. For regular SAXS experiments, the DM is moved out of the beam path and has no obstacle to the beam. The optical layout of this beamline is shown in Figure 1.

# Coordinate System

Laboratory Coordinate is listed below:

- The beam propagation coincides with the Y axis.

- Z axis shows the vertical direction.

- X axis is perpendicular to both Y and Z axes.

- Pitch(θ): rotation  around X axis.

- Roll(χ):  rotation around Y axis.

- Yaw(φ): rotation around the Z axis.

The definitions of the coordinate axes fixed to the mirror are listed below:

- X is the sagittal direction, i.e. along the mirror surface transverse to the reflection/diffraction plane.

- Y is in the meridional direction i.e. along the mirror surface within the reflection/diffraction plane.

- Z points away from, and is normal to the mirror, i.e. along the mirror surface normal.

- Pitch(θg): rotation around X axis(grazing angle).

- Roll(χ):  rotation around Y axis.

- Yaw(φ): rotation around the Z axis.

# Specifications

## Specifications of the chamber

The conceptual drawing of the deflection mirror system is shown in figure 3 and the specifications and the relevant informations are listed below:

- Energy range:9keV-15keV.

- Beam height at the center of the mirror: 1384.0mm.

- The vertical distance from the surface of the optical table to the beam is 25.0mm.

- The distance from the source to the mirror center is 44.0m.

- Beam cross section at 44m:6.5mm×1.0mm(H×V).

- Grazing angle of mirror: 0.3°.

- The mirror is facing downward(-Z). Its geometry is shown in Figure 2.

- A rectangular shape is preferred for the chamber. Its length is limited to 500mm.

- All motorized adjustments should have absolute encoders. The type of encoder and its protocol must be approved by NSRRC.

- Adjustments for the deflection mirror: see table 1. Stewart platform(Hexapod) is preferred including its controller and software. The Stewart platform should be compatible with EPICS.

- The vacuum chamber shall be made of 304 stainless steel and cleaned in line with the vacuum requirements.

- Two sets of fiducial mounts for laser tracker.

- The whole system can be manually adjusted for alignment via screws.

- The ion pump should be independently supported to prevent transferring its load to the chamber.

- A screen is used to monitor the input beam. Two limit switches are used to define the two extreme positions of movement. The types of the limit switches and the position sensors should be approved by NSRRC.

- Ultimate pressure: better than 1×10-6 Torr.

Figure 1: Conceptual optical layout of TPS 14A beamline.

Leak rate: better than 10-10 Torr.L/s(Helium).

Figure 2: Geometry of deflection mirror. The weight is ~405g..

Figure 3: Conceptual drawing(side view) for the DM chamber system.

Table 1: Adjustments of the deflection mirror.

| Adjustment | Type | Range | Resolution | Stability(2 hours) |
| --- | --- | --- | --- | --- |
| X | Motorized | ±10mm | <5μm | <1μm |
| Z | Motorized | ±5mm | <5μm | <1μm |
| Ptich | Motorized | ±1° | <20μrad | <10μrad |
| Roll | Motorized | ±1° | <20μrad | <10μrad |
| Yaw | Motorized | ±1° | <20μrad | <10μrad |

Table 2: Specifications of the flanges for the chamber.

| Port | Quantity | Size | Position |
| --- | --- | --- | --- |
| Entrance | 1 | 40CF | centered at the input beam |
| Exit | 1 | 40CF | centered at the output beam |
| Ion pump | 1 | TBD/Rotatable |  |
| View port | 1 | 63CF | for motoring the input beam |
| Electrical | TBD | 63CF |  |
| Guage | 1 | 40CF |  |
| Pumping port | 1 | 40CF |  |

Table 3: Specifications of the chamber's coarse adjustments

| Adjustment | Range | Resolution |
| --- | --- | --- |
| X | ±10mm | 5μm |
| Z | ±10mm | 5μm |
| Roll | ±2° | 0.001° |
| Pitch | ±2° | 0.001° |

## Control

NSRRC will provide the motion controller for the adjustments. All motors/actuators should be compatible with the controller provided by NSRRC. All controllers provided by the vendor should be compatible with EPICS.

## Specifications of the deflection mirror(for reference)

The specifications of the deflection mirror are listed below(See Figure 2.):

- Position:44.0m from the source

- Type:plane mirror

- Facing: downward

- Substrate: Silicon.

- Useful area.:250mmx50mm(LxW).

- Coating:Pt

- Curvature:R>50km, ⍴>10km

- Focal point:infinite

- Object distance:infinite

- Image distance:infinite

- Grazing angle:0.25°(~4.36mrad)

- Slope error:<1.0μrad

- Cooling:none

- Roughness: <0.3nm.

- Ultra-high Vacuum compatible.

- Alignment mark

- Marks for coatings

# Inspection and Acceptance Criteria

## General terms

- In order to meet these specifications, NSRRC reserves the right to inspect and require alterations of the completed design.

- Agreement between NSRRC and the vendor on the entire design drawings must occur before any actual fabrication process begins.

- Unless otherwise specified in the contract or order, the vendor is responsible for the performance of all inspection requirements as specified herein.

- Except as otherwise specified, the vendor may utilize his own facilities or any commercial laboratory acceptable to NSRRC.

- NSRRC should be informed well first before acceptance tests are being performed so that at least one of NSRRC staff can be present during the tests.

- A written report of the tests shall be provided to NSRRC to verify compliance with the specifications. Test methods, criteria and measuring equipment should be shown in the test report.

- The vendors shall foresee the insurance of the equipment during the tests, and the transport to the metrology laboratory.

- The components supplied by NSRRC useful to do the tests will be loaned to the vendor premises on request. However, this does not release the vendor from the responsibility of proving the performance of the components.

## Factory acceptance tests

Factory acceptance tests report shall provide the following test items:

- Exterior appearance check

  - Configuration should accord with the approval drawing.

  - No flaw, dent, scratch and fouling that affect the equipment in use and/or in function.

  - The prescribed visual range is covered in each viewport.

  - No loose bolt or screw.

  - No sharp edges and corners.

  - Cables should be correctly labeled.

- Dimensional check.

- Helium leakage test.

- Movement test

  - Each limit switch should perform properly.

  - Stroke/range, resolution, repeatability and stability should be tested for each motorized movement.

  - Stroke/range should be tested for each manual movement.

  - Each encoder should operate normally.

- Electrical circuit test.

- Ultimate vacuum pressure test.

- Fiducial point check.

## On-site Performance and Functional Tests

After on-site installation, the performance and function should be tested by the vendor. The report shall provide the following test items:

• Helium leakage test

• Movement test

◦ Stroke/range and resolution should be tested for each motorized movement.

◦ Each encoder should operate normally.

The following tests will be performed by NSRRC after on-site installation:

- Helium leakage test

- Movement test

  - Stroke/range and resolution should be tested for each motorized movement.

  - Each encoder should operate normally.

## User manual

Vendor should provide a user manual and the following information should be included

- Safety

- System description

- Installation

- Operation

- Maintenance, repair and disposal

- Troubleshooting

- Drawings

# Packing and Shipping

The finished components must be thoroughly clean and free from any organic materials or cutting fluid. The equipment must be packed in such a way that foreign materials cannot penetrate inside. The packaging methods must be subject to approval by NSRRC.

Package for shipment must ensure that the equipment is insulated from any severe shock. Markings indicating the fragile nature should be on the outside wall of the container.

The quoted item(s) shall be completed, packaged, delivered DPU NSRRC unloaded (Incoterms 2020) by the vendor.

The shipping address is:

National Synchrotron Radiation Research Center

101 Hsin-Ann Rd., Hsinchu 30076, Taiwan

Attention: Dr. Chi-Yi Huang

Besides the shipping address, the following is to be displayed clearly on the outside of the packaging:

- NSRRC contract number

- Weight of the loaded package

- Support points for transport and lifting

- Tilt indicators

- Shock indicator

- Dimensions of the box

# Warranty

- The vendor warrants that all Contract Product(s) delivered by the vendor to NSRRC shall be of new manufacture, free from defects in material, workmanship.

- The warranty will remain in force for 12 months from installation and NSRRC's acceptance.