Technical Specifications :

Double Multilayer Monochromator at SP12U

National Synchrotron Radiation Research Center

Hsinchu, Taiwan R.O.C.

(OCT, 2025)

CONTENTS

1	SCOPE OF THE CONTRACT

2	REQUIREMENTS

3	INSPECTION AND ACCEPTANCE CRITERIA

4	DELIVERY

- SCOPE OF THE CONTRACT

The scope of this contract is the fabrication of double multilayer monochromator (DMM) mirrors. Each mirror shall be based on a silicon crystal substrate, with the top surface polished to achieve a slope error of less than 0.5 µrad. The substrate dimensions must be 120 mm in length, 50 mm in width, and 30 mm in thickness. Each mirror must feature four distinct stripes on the top surface, with varying multilayer periods and layer counts, designed to reflect X-rays at approximately 8, 11, 13, and 15 keV, respectively, at an incident angle of 1.48°. The multilayers are expected to consist of Mo (or Ni) and B₄C, with 100 to 400 layers and a periodicity ranging from 1.6 to 3.0 nm. Three sets of mirrors are required: two for operational use and one as a spare.

Fig.1 Schematic diagram of DMM mirrors

- Requirements

The requirements are summarized in the following follows.

| Shape | Flat | Remarks |
| --- | --- | --- |
| Dimension | 120 (L) × 50 (W) × 30 (T) mm3 |  |
| Substrate | Czochralski (or floating-zone) pure Si (100) surface | Resistivity ≥ 5.0 (Ω・cm) |
| Useful area (polished area) | ≥100 (L) × 25 (W) mm2 |  |
| Tangential Slope error | ≤0.5 µrad RMS |  |
| Sagittal Slope error | ≤1 µrad RMS |  |
| Surface roughness | ≤ 0.2 nm | Before coating |
| Curvature in tangential | ≥ 50 km |  |
| Curvature in sagittal | ≥ 10 km |  |
| Coating | Mo / B4C, 4 lanes | See below |
| Coating area | ≥100 (L) × 4 (W) mm2 | On Cr binder |
| Coating margin | 2 mm |  |
| Quantity | 3 | One is spare |

Table 1 Fundamental specification of DMM substrate

|  | Lane I | Lane II | Lane III | Lane IV |  |
| --- | --- | --- | --- | --- | --- |
| Energy (at θB = 1.48°) | ~ 8.0 keV | ~ 10.9 keV | ~ 13.3 keV | ~ 15.0 keV |  |
| D-spacing | 3.0 ± 0.05 nm | 2.2 ± 0.05 nm | 1.8 ± 0.05 nm | 1.6 ± 0.05 nm |  |
| Gamma : Z(B4C)/D | 0.5 ± 0.03 | 0.5 ± 0.03 | 0.5 ± 0.05 | 0.5 ± 0.05 |  |
| Number of periods | ≥ 120 | ≥ 275 | ≥ 375 | ≥ 400 |  |
| Materials | Mo (or Ni), B4C | Mo, B4C | Mo, B4C | Mo, B4C |  |

Table 2 Fundamental specification of multilayer coating

- INSPECTION AND ACCEPTANCE CRITERIA

  - Visible inspection

Any visible damages, such as clacks or dips of > 0.1 mm, are not accepted for the top surface. Meanwhile, those of > 1 mm are not accepted for the other surfaces.

  - Dimensions

The length, width and thickness of substrates should be accurate within 0.2 mm while the length and width of the coating within 0.1 mm (Table 1)

  - Factory acceptance test (FAT) report

The supplier should provide NSRRC with the factory acceptance test (FAT) report. The measured parameters must be compliant with those specified in Tables 1 and 2.

  - Material report

The supplier should provide NSRRC with the material report for the substrate. The substrate should be made of floating-zone pure Si (Table 1).

- DELIVERY

The products must be delivered in 360 calendar days by the date of awarding contract. The supplier is responsible for the delivery of all the components, including the payment of the import/consumption tax. The address for the delivery is BL12XU at SPring-8, 1-1-1 Kouto, Sayo-cho, Sayo-gun, Hyogo 679-5198, Japan (DDP, Hyogo, JP, INCOTERM2020)