Specifications of advanced Kirkpatrick–Baez mirrors at SPring-8 II BL12XU Taiwan Beamline

National Synchrotron Radiation Research Center

101 Hsin-Ann Rd., Hsinchu 30076, Taiwan

2026年5月11日

The National Synchrotron Radiation Research Center (hereinafter referred to as "NSRRC")

shall be pleased to receive full details of your quotation in accordance with the following terms and conditions:

Overviews

The document is for the bid of advanced Kirkpatrick–Baez mirrors at SPring-8-II BL12XU Taiwan Beamline, Hyogo 679-5198 Japan.

Scope of the contract

The vendor should provide:

- One set of advanced Kirkpatrick–Baez mirrors including one vertical focusing mirror and one horizontal mirror.

- Factory test report.

Coordinate System

Laboratory Coordinate is listed below:

- The beam propagation coincides with the Y axis.

- Z axis shows the vertical direction.

- X axis is perpendicular to both Y and Z axes.

- Pitch(θ): rotation  around X axis.

- Roll(χ):  rotation around Y axis.

- Yaw(φ): rotation around the Z axis.

The definitions of the coordinate axes fixed to the mirror are listed below:

- X is the sagittal direction, i.e. along the mirror surface transverse to the reflection/diffraction plane.

- Y is in the meridional direction i.e. along the mirror surface within the reflection/diffraction plane.

- Z points away from, and is normal to the mirror, i.e. along the mirror surface normal.

- Pitch(θg): rotation around X axis(grazing angle).

- Roll(χ):  rotation around Y axis.

- Yaw(φ): rotation around the Z axis.

Specifications

Specifications of advanced Kirkpatrick–Baez(AKB) mirrors:

The AKB is composed of one vertical focusing mirror(VFM) and one horizontal focusing mirror(HFM). The VFM has two reflecting surfaces. The upstream one is a plane-elliptical mirror and the downstream one is a plane-hyperbolic mirror. The HFM has the same configuration.

The specifications of the VFM are listed below:

- Name: VFM-ellipse

- Position: 72.0647m from the source.

- Shape of reflecting surface: plane-elliptical.

- Substrate: CZ-Si(100).

- Facing: upward (+Z).

- Optical aperture:>4.0mm×135mm(W×L).

- Grazing angle: 3.35mrad.

- Object distance: 72064.7mm.

- Image distance: 1680mm.

- Meridional slope error: ≤ 0.1µrad(rms).

- Sagittal slope error: ≤10µrad (rms).

- Roughness: <0.3nm at 50X.

- Coating: 50 nm±10 nm Pt with Cr buffer.

- Alignment mark.

- Name: VFM-hyperbolic

- Position: 72.22m from the source.

- Shape of reflecting surface: plane-hyperbolic.

- Substrate: CZ-Si(100).

- Facing: upward (+Z).

- Optical aperture:>4.0mm×132.2mm(W×L).

- Grazing angle: 3.11mrad.

- Object distance: 1524.7mm.

- Image distance: 780mm.

- Meridional slope error: ≤ 0.1µrad(rms).

- Sagittal slope error: ≤10µrad (rms).

- Roughness: <0.3nm at 50X.

- Coating: 50 nm±10 nm Pt with Cr buffer.

- Alignment mark.

The total dimension of the VFM is 303.3mmx50mmx50mm(LxWxT). The tolerance of the length is +/- 0.5mm. The tolerance of the other dimension is +/- 0.3mm.

The specifications of the HFM are listed below:

- Name: HFM-ellipse

- Position: 72.379m from the source.

- Shape of reflecting surface: plane-elliptical.

- Substrate: CZ-Si(100).

- Facing: horizontal (X).

- Optical aperture:>4.0mm×135mm(W×L).

- Grazing angle: 3.35mrad.

- Object distance: 72379mm.

- Image distance: 1058mm.

- Meridional slope error: ≤ 0.1µrad(rms).

- Sagittal slope error: ≤10µrad (rms).

- Roughness: <0.3nm at 50X.

- Coating: 50 nm±10 nm Pt with Cr buffer.

- Alignment mark.

- Name: HFM-hyperbolic

- Position: 72.535m from the source.

- Shape of reflecting surface: plane-hyperbolic.

- Substrate: CZ-Si(100).

- Facing: horizontal(X).

- Optical aperture:>4.0mm×126mm(W×L).

- Grazing angle: 3.06mrad.

- Object distance: 902mm.

- Image distance: 465mm.

- Meridional slope error: ≤ 0.1µrad(rms).

- Sagittal slope error: ≤10µrad (rms).

- Roughness: <0.3nm at 50X.

- Coating: 50 nm±10 nm Pt with Cr buffer.

- Alignment mark.

The total dimension of the HFM is 298.3mmx50mmx50mm(LxWxT). The tolerance of the length is +/- 0.5mm. The tolerance of the other dimension is +/- 0.3mm.

Inspection and Acceptance Criteria

General term

- In order to meet these specifications, NSRRC reserves the right to inspect and require alterations of the completed design.

- Agreement between NSRRC and the vendor on the entire design drawings must occur before any actual fabrication process begins.

- Unless otherwise specified in the contract or order, the vendor is responsible for the performance of all inspection requirements as specified herein.

- Except as otherwise specified, the vendor may utilize his own facilities or any commercial laboratory acceptable to NSRRC.

- A written report of the tests shall be provided to NSRRC to verify compliance with the specifications. Test methods, criteria and measuring equipment should be shown in the test report.

- NSRRC shall be informed well first before the factory acceptance tests so that at least one of NSRRC staff can be present during the tests.

- The vendors shall foresee the insurance of the equipment during the tests, and the transport to the metrology laboratory.

Factory acceptance tests

Factory acceptance tests report shall show the following items:

- Approval by NSRRC staff: print name of NSRRC staff attending the tests and his/her signature shall be shown in the report.

- Tester of the vendor: print name and his/her signature

- Date

- Place

- Test methods

- The brand and model of the test equipment.

- Exterior appearance check

  - Configuration should accord with the approval drawing.

  - No flaw, dent, scratch and fouling that affect the equipment in use and/or in function.

  - No sharp edges and corners.

- Dimensional check.

- Alignment mark.

- Roughness

- Curvature and slope error

Delivery

Maximum number of partial deliveries of product is 2 times.

Payment Terms

Pay according to the price listed in “Price List of Tender” after inspection and acceptance of each mirror.

Packing and Shipping

The finished components must be thoroughly clean and free from any organic materials or cutting fluid. The equipment must be packed in such a way that foreign materials cannot penetrate inside. The packaging methods must be subject to approval by NSRRC.

Package for shipment must ensure that the equipment is insulated from any severe shock. Markings indicating the fragile nature should be on the outside wall of the container.

The supplier must be responsible for the delivery of all the components, including the payment of the import/consumption tax. The address for the delivery is BL12XU at SPring-8, 1-1-1 Kouto, Sayo-cho, Sayo-gun, Hyogo 679-5198, Japan (DDP, Hyogo, JP, INCOTERM2020)  Attention: Dr. Nozomu Hiraoka, Consignee: Japan Synchrotron Radiation Research Institute.

Besides the shipping address, the following is to be displayed clearly on the outside of the packaging:

- NSRRC contract number

- Weight of the loaded package

- Support points for transport and lifting

- Tilt indicators

- Shock indicator

Warranty

- The vendor warrants that all Contract Product(s) delivered by the vendor to NSRRC shall be of new manufacture, free from defects in material, workmanship.

- The warranty will remain in force for 12 months at least from installation and NSRRC's acceptance.