Technical Specifications :

The vacuum chamber and adjustment systems for advanced KB mirrors at SP12U

National Synchrotron Radiation Research Center

Hsinchu, Taiwan R.O.C.

(MAR, 2026)

CONTENTS

- SCOPE OF THE ORDER

- EXISITING INSTRUMENTATION AND CURRENT STATUS

- REQUIREMENTS

- INSPECTION AND ACCEPTANCE CRITERIA

- DELIVERY

- RESPONSIBILITIES OF THE SUPPLIER

- SCOPE OF THE ORDER

The scope of the present order includes the design and fabrication of a vacuum chamber and mirror adjustment systems for Kirkpatrick–Baez (KB) mirrors or advanced KB (AKB) mirrors. The mirror dimensions are common between KB and AKB; each pair consists of mirrors with a width of 50 mm, a height of 50 mm, and a length of ~300 mm. The chamber must be compatible with ultra-high vacuum (UHV). The mirror positions and angles must be adjustable by remote motors within the UHV environment. The chamber and adjustment systems must be carefully designed because the mirrors will be positioned close to the sample, mounted at the top of the positioning tower, with a working distance of approximately 400 mm during operation. When not in use, the mirrors must be retracted to a distance of approximately 1.9 m from the sample to provide a larger working space. The vendor is required to install the chamber and adjustment system into the IXS end-station in SP12U. A movable granite base will be installed on linear guide rails by NSRRC. The vendor is requested to fix the chamber and adjustment systems on the granite. The pump groups, the vacuum gauges, and the motor drivers are prepared with NSRRC. The pumping ports and motor cable connectors shall be prepared by the vendor.

- EXISITING INSTRUMENTATION AND CURRENT STATUS

The mirror dimensions parameters are shown below. The new vacuum chamber must accommodate the V-(A)KB, H-(A)KB mirrors and their stages. They will be installed between the sample tower located at the position from 73m from the light source and the circular rail mounting an X-ray spectrometer equipped with multiple analyzers, existing at SP12U beamline at SPring-8. The radius of the circular rail is 2.5 m. When KB or AKB mirrors are in operation, the spectrometer positions at a scattering angle in 0-115 deg. When not in operation, the spectrometer is in 0-150 deg. The chambers and stages accommodating (A)KB mirrors must not touch the sample tower and the spectrometer. Another critical issue is that the beam height is different when the (A)KB is operational and not operational: 1430 mm from the floor in farmer while 1495 mm in latter. The both beams must pass through the chamber.

|  | V-(A)KB | H-(A)KB |  |
| --- | --- | --- | --- |
| Length [mm] | 300±5 | 300±5 |  |
| Width [mm] | 50±0.2 | 50±0.2 |  |
| Height [mm] | 50±0.2 | 50±0.2 |  |
| Position from source [m]+ | 72.06±0.1 | 72.38±0.1 |  |
| Incident angle [mrad] | 0 - 8 mrad | 0 - 8 mrad |  |
| Note+: Position of the center of the elliptical surface of AKB from the source |  |  |  |

Table 1. Mirror parameters

- REQUIREMENTS

  - Dimension of the vacuum chamber

The vacuum chamber accommodates the V-(A)KB, H-(A)KB mirrors and their stages with motors. The whole set will be mounted on a heavy-duty motorized XYZ stages. The working distance of the mirror downstream is 400 mm, that is, the downstream edge of the H-(A)KB is positioned 400 mm away from the sample position. However, the real working space will be determined by the flange for the beam exit. We request the real working distance of > 250 mm. The dimension of the sample positioning tower is shown in Fig 1. The vendor is requested careful discussions with NSRRC to avoid the interference between the (A)KB chamber and the sample tower. In addition, when (A)KB is not operational, the whole system is retracted 1.2-m upstream and the x-ray spectrometer arm is rotated up to 150 deg. The vendor also should take care to avoid the touch between the mirror chamber and the spectrometer.

  - In-vacuum stages

Table 2 shows the stages required to be in the vacuum chambers and their specifications. Four manual or remote axis are required for each mirror.

  - Out-vacuum stages

Three remote axis (XYZ) plus one manual axis (θz) are required (Table 3). They need to be heavy duty as > 200 kg.

  - Auto-collimator

An auto-collimator is required to ensure the surfaces on VKB and HKB are orthogonal to each other within an accuracy 1 mrad. A common way may be to use a laser to find reflection from the surfaces. The auto-collimator may be attached onto the top of the chamber.

  - Front door

A front door should be prepared to access to the mirrors for installation or deinstallation. The opening should be larger than 500 x 500 mm2. The hinges should be on upstream side. Two type of sealing is requested to prepare: One is the O-ring sealing while the other the metal wire sealing.

  - Flanges

Two ICF 152 flanges are required for the beam entrance and exit. The flanges with Be windows are prepared by NSRRC. Another ICF203 and ICF253 flanges, plus six ICF70 flanges are required for vacuum pumps and feed-though on the back, respectively. Feed-through should be prepared by the vendor while pumps are prepared by NSRRC.

  - Motors and connectors

Five-phase motors must be used for the remote axes. They should be terminated with 8-pin circular connectors (Fig 3) based on SPring-8 standard.

  - Vacuum

It must be ensured that the chamber keeps the vacuum of < 2.0E-4 Pa during the operation with pumping (turbo molecular and/or ion pump).

In-vacuum, V-(A)KB (Upstream)

| Name of axis | Type | Range | Resolution |
| --- | --- | --- | --- |
| V-θz | Yaw / manual | ±>1 deg | --- |
| V-θy | Roll / manual | ±>0.5 deg | --- |
| V-θx* | Pitch / remote | ±> 20 mrad | ≤0.01mrad |
| V-Y | Vertical linear / remote | ±> 12.5 mm | ≤10μm |
| Note*: The center of the rotation should be at the center of the AKB elliptical surface |  |  |  |

In-vacuum, H-(A)KB (Downstream)

| Name of axis | Type | Range | Resolution |
| --- | --- | --- | --- |
| H-θz** | Yaw / remote | ±>20 mrad | ≤0.01mrad |
| H-Z | Vertical, linear / remote | ±> 10.0 mm | ≤10μm |
| H-θx | Pitch / manual | ±> 1 deg | --- |
| H-X | Horizontal, linear / remote | ±> 12.5 mm | ≤10μm |
| Note**: The center of the rotation should be at the center of the AKB elliptical surface |  |  |  |

Table 2. In-vacuum remote / manual axes

Out-vacuum, Common

| Name of axis | Type | Range | Resolution |
| --- | --- | --- | --- |
| VH-Z | Vertical, Linear, remote | ±>15 mm | ≤10μm |
| VH-X | Horizontal, Linear, remote | ±>15 mm | ≤10μm |
| VH-Y | Horizontal, Linear, remote | ±>15 mm | ≤10μm |
| V-θz | Yaw/ manual | ±> 1 deg | --- |

Table 3. Out-vacuum remote / manual axes

- INSPECTION AND ACCEPTANCE CRITERIA

  - Visible check

No visible damage is ensured. Dips or scratches deeper than 1.0 mm and longer than 5 mm are not permitted.

  - Factory acceptance report

The supplier must provide NSRRC with the factory acceptance test reports. The travel ranges and resolutions must be ensured for all of the manual / remote axes described in (3.2, 3.3). The autocollimator should have an angular accuracy of 1 mrad (3.4) in measurement.

  - Drawings

The supplier must provide NSRRC with the drawings. Two types of sealing on the front door (3.5) and the flanges specified in (3.6) must be ensured. A wire diagram for the motor cables and connectors also must be provided (3.7).

  - Vacuum test / Leak test

A vacuum test must be performed. A vacuum of <2.0E-04 Pa with and without the Z-motion (3.8). Alternatively, a leak test with He gas may be performed. A leak rate < 2.0E-10 Pa m3/s must be achieved.

  - On-site function test report

The supplier is responsible for the on-site function test, which will be supervised by NSRRC. Smooth motions and the proper functioning of limit switches must be ensured for all motorized axes.

- DELIVERY

The supplier is responsible for the delivery of all the components, including the payment of the importing/consuming tax. The address for the delivery is BL12XU at SPring-8, 1-1-1 Kouto, Sayo-cho, Sayo-gun, Hyogo 679-5198, Japan (DDP, Hyogo, JP, INCOTERM2020)

- RESPONSIBILITIES OF THE SUPPLIER

  - The machining drawings should be provided. Several discussions are required with NSRRC before the details are finalized. The supplier should provide NSRRC with the drawings within 60 calendar days after awarding the contract. The supplier should deliver the products within 360 calendar days after receiving the notification of drawing confirmation from NSRRC. The installation and function test should be completed within 60 calendar days after the notification from NSRRC.

  - The manufacturing and the transportation of the required products to the SPring-8 site.

  - Factory acceptance test (FAT) reports

  - On-site function test report

  - Vacuum/Leak test

Fig. 1 Geometry of AKB chamber and the sample tower. AKB and the chamber/adjustment system will be mounted on the granite block, which will be translated by 1.2 m along the linear guide rails. The vendor is required to design the chamber/adjustment system positioning at the right position without a touch. The granite and the linear guide rails are prepared by NSRRC.

Fig. 2 Geometry of AKB chamber and the spectrometer arm. The arm will be rotated up to 115 deg when the mirror chamber is the closest position (Pos.I) while up to 150 deg when the mirror is farthest (Pos.II). The vendor is required to design the chamber/adjustment system positioning at those positions without a touch.

Fig.3 Eight-pin connector; A,B,C,D,H: Power (5 phase) E: LS-COMMON F: CW-LS G: CCW-LS