National Synchrotron Radiation Research Center

Procurement and Specification for

Cylindrically Bent Crystal Analyzer (SRPO-1130357)

The vendor is required to supply three cylindrically bent crystal analyzers (CBCAs), specifically for the Si(531), Ge(440), and Ge(444) orientations.

- Specification:

- Cylindrically bent crystal analyzer (CBCA)

- Orientation: Si(531), Ge(440) and Ge(444)

- Quantity: One analyzer per specified orientation

- Material: Silicon for Si(531) orientation; Germanium for Ge(440) and Ge(444) orientations

- Radius of Curvature: 500 mm

- Crystal Thickness: 300 ±15 um

- Surface Finish: Polished surface

- Surface Defects: Less than 1%

- Wafer Size: 100 x 50 mm2

- Bottom Substrate Size: 101.6 x 50.8 mm2

- Aluminum Substrate Base: 101.6 x 50.8 x 5 mm³ (length x width x height) with 3-M4 mounting holes

- Warranty:

- The vendor is required to provide a one-year warranty on the product.

- Inspection & Acceptance:

- Item Confirmation

- Three cylindrically bent crystal analyzers (CBCAs) for Si(531), Ge(440) and Ge(444) orientations

- The vendor shall provide the following documents as specified below:

- Certificate of Manufacture

- Factory Acceptance Test (FAT) report