Technical Specifications:

Ion pump controllers and vacuum gauge controllers for Front-ends in SP12U

National Synchrotron Radiation Research Center

Hsinchu, Taiwan R.O.C.

(APR, 2026)

CONTENTS

1	SCOPE OF THE CONTRACT

2	REQUIREMENTS

3	INSPECTION AND ACCEPTANCE CRITERIA

4	DELIVERY

- SCOPE OF THE ORDER

The scope of this contract is to supply vacuum gauges and controllers for the ion pumps installed in the SP12U front-end. The existing ion pumps will continue to be used, while the gauges and controllers will be replaced with new models. The new gauges and controllers must be fully compatible with the existing ion pumps.

- Requirements

|  | Model | Quantity |
| --- | --- | --- |
| Vacuum gauge controllers | Canon-Anelva M923-DD | 9 |
| Ion pump controllers | Canon-Anelva P-521NP | 9 |

The vacuum gauge controllers and ion pump controllers must be compatible the existing ion pumps, Canon-Anelva Noble pumps 912-7020 (110L/s), 912-7160 (60L/s), and 7130 (30L/s). The required specifications are follows:

  - Vacuum gauge controller

- Pressure range: 1.0E-8 Pa – 1.0E-02 Pa

- Filament protection function available

- Degassing mode function available

- Switch contacts: 2 setpoints in pressure (or current) function available

- Remote / Local mode switching function available

  - Ion pump controller

- Output voltage: DC-5.2kV±10% / DC-7.0kV±10% (switchable)

- INSPECTION AND ACCEPTANCE CRITERIA

  - Visible inspection

Any visible damages, such as clacks or dips of > 2 mm, are not accepted.

  - Factory acceptance test (FAT) report

The supplier should provide NSRRC with the factory acceptance test (FAT) reports. For the vacuum gauge controller, the display range and other functions specified in 2.1 shall be verified to be available. For the ion pump controller, the output voltages shall be examined for each unit and verified to fall within the range specified in 2.2.

- DELIVERY

The supplier is responsible for the delivery of all the components, including the payment of the import/consumption tax. The address for the delivery is BL12XU at SPring-8, 1-1-1 Kouto, Sayo-cho, Sayo-gun, Hyogo 679-5198, Japan (DDP, Hyogo, JP, INCOTERM2020)