# Specification of High-Precision Vacuum Viewport

## 1. Introduction

This document specifies the requirements for a high-precision vacuum viewport to be installed on the in-situ LTP vacuum chamber for optical observation, laser/beam transmission, and related measurement applications. The viewport shall be suitable for operation in high vacuum or ultra-high vacuum environments and shall provide the required optical performance for the intended system.

## 2. Scope and Applicable Documents

This specification establishes the requirements for the fabrication, inspection, testing, packaging, and delivery of the high-precision vacuum viewport assembly.

The supplier shall provide the following documents, where applicable:

- Product specification sheet or catalog

- Dimensional information or drawing

- Material description for the window

- Coating specification

- Inspection and acceptance documents

- Other supporting technical documents, if required

Additional technical information may be requested from the awarded supplier if necessary.

## 3. Requirements

### 3.1 General

The vacuum viewport shall be supplied as a complete assembly suitable for installation on the existing vacuum chamber system.

### 3.2 Quantity

Quantity: 2 units

### 3.3 Operating Conditions

- Vacuum environment: High Vacuum / Ultra-High Vacuum (UHV)

- Bakeout temperature: 200°C or higher

- Operating temperature: Ambient condition unless otherwise specified

### 3.4 Geometrical and Optical Parameters

- Item: High-Precision Vacuum Viewport

- Flange type: 8.00-inch ConFlat (8.00CF)

- Window material: UV-grade or DUV-grade fused silica

- View diameter / clear aperture: 5.38 inches or larger

- Optical flatness: λ/10 @ 632.8 nm or better

- Surface quality: 20/10 scratch-dig or better

- Anti-reflective coating: Single-wavelength AR coating at 635 nm

- Coating arrangement: Single-side coating on the air side only

- Reflectance: R < 0.25% @ 635 nm or better

- Vacuum compatibility: Suitable for UHV

- Bakeout capability: 200°C or higher

- Sealing / joining method: The fused silica window shall be hermetically sealed to the metal flange by brazing.

## 4. Quality Assurance

The supplier shall be responsible for all inspection requirements specified herein. The supplier may use in-house inspection facilities or other qualified facilities acceptable to the purchaser. The purchaser reserves the right to perform inspections necessary to ensure compliance with this specification.

### Table 1. Specification of High-Precision Vacuum Viewport

| Specification Item | Requirement |
| --- | --- |
| Item | High-Precision Vacuum Viewport |
| Quantity | 2 units |
| Flange type | 8.00-inch ConFlat (8.00CF) |
| Window material | UV-grade or DUV-grade fused silica |
| View diameter / clear aperture | 5.38 inches or larger |
| Optical flatness | λ/10 @ 632.8 nm or better |
| Surface quality | 20/10 scratch-dig or better |
| AR coating | Single-wavelength AR coating at 635 nm |
| Coating arrangement | Single-side coating on the air side only |
| Reflectance | R < 0.25% @ 635 nm or better |
| Vacuum compatibility | UHV |
| Bake out temperature | 200°C or higher |

### Table 2. NSRRC and FAT Inspection Items

| Parameter | Requirement | NSRRC Inspection | FAT* measuring items |
| --- | --- | --- | --- |
| Viewport condition | No visible damage | Visual inspection | FAT |
| Flange type | 8.00CF | NSRRC inspection | FAT |
| Quantity | 2 units | NSRRC inspection | FAT |
| Window material | UV-grade or DUV-grade fused silica | Document review | FAT |
| View diameter / clear aperture | 5.38 inches or larger | NSRRC inspection | FAT |
| Optical flatness | λ/10 @ 632.8 nm or better | Document review | FAT |
| Surface quality | 20/10 scratch-dig or better | Document review | FAT |
| Coating arrangement | Single-side coating on the air side only | Document review | FAT |
| AR coating | Single-wavelength AR coating at 635 nm | Document review | FAT |
| Reflectance | R < 0.25% @ 635 nm or better | Document review | FAT |
| Leak rate | Shall be less than 1.013 × 10^-9 mbar·L/s He | NSRRC leak test | FAT |
| Bake out temperature | 200°C or higher | Document review | FAT |

*FAT: Factory Acceptance Test

## 5. Preparation for Delivery

### 5.1 Cleaning

The finished product shall be thoroughly cleaned and free from dust, oil, grease, machining residue, and other contaminants that may affect vacuum performance.

### 5.2 Delivery

The products shall be delivered to the designated site in accordance with the purchase order or contract requirements.

### 5.3 Packaging

Each viewport shall be packaged appropriately to prevent dust contamination, moisture ingress, and mechanical damage during transportation and storage.

### 5.4 Packing

Packing for shipment shall be sufficient to protect the product from shock, vibration, and rough handling during transportation.

## 6. Payment Terms

Payment shall be made in accordance with the NSRRC procurement and payment procedures after delivery and acceptance of the goods.