Specification of Vacuum Compatible Positioners

SRPO-1150209

- Condenser Stage

  - Quantity: 1

  - Environment: high vacuum (1e-6 mbar)

  - Material: stainless steel components

  - Coordinate system:

  - Dimension: 90±15mm x 130±15mm x 140±15mm (length x width x height)

  - Positioner:

    - X1: >120 mm travel range, integrated optical encoder

    - X2: >120 mm travel range, integrated optical encoder

    - Y1: > 80 mm travel range, integrated optical encoder

    - Y2: > 80 mm travel range, integrated optical encoder

    - Rz: 360 ° travel range, integrated optical encoder

    - Gx: 5 ° travel range

    - Z: >15 mm travel range, integrated optical encoder

- System cabling:

    - X: 2x 80 cm cable, Teflon-stranded wires with braid, each 1x circular 14-pin LEMO plug, male

    - Y: 2x 80 cm cable, Teflon-stranded wires with braid, each 1x circular 14-pin LEMO plug, male

    - Rz: 2x 80 cm cable, Teflon-stranded wires with braid, each 1x circular 14-pin LEMO plug, male

    - Gx: 2x 80 cm cable, Teflon-stranded wires with braid, each 1x circular 14-pin LEMO plug, male

    - Z: 1x 80 cm cable, each 1x circular 14-pin LEMO plug, male

  - Sensor resolution: < 1um

  - Moving speed: >10 mm/s

  - Mechanical components: These mechanical components can be systematically integrated to construct a high-precision, multi-axis positioning stage. Standard basis adapter (base), planar standard adapter (x-y), customized planar adapter (y-Rz), planar standard adapter (Rz-Gx), rectangular standard adapter (Gx-Z), rectangular standard adapter (Z-sample), single constant force spring (Z, lift force: 3N )

  - System assembly (assembly, final test)

  - Factory Acceptance Test (functional test) shall include verification of travel range, resolution, material.

- Phase ring stage

  - Quantity: 1

  - Environment: high vacuum (1e-6 mbar)

  - Material: stainless steel components

  - Coordinate system:

  - Dimension: 35±5mm x 20±5mm x 35±5mm (length x width x height)

  - Positioner:

    - X: > 10 mm travel range, integrated optical encoder

    - Y: > 10 mm travel range, integrated optical encoder

    - Z: > 10 mm travel range, integrated optical encoder

  - Sensor resolution: < 1um

  - Moving speed: >10 mm/s

- System cabling: 3x 80 cm cable, Teflon-stranded wires with braid, common shielded braid, 1x circular 32-pin LEMO plug, male

  - Mechanical components: These mechanical components can be systematically integrated to construct a high-precision, multi-axis positioning stage. Standard basis adapter (base), planar standard adapter (x-y), rectangular standard adapter (Z-sample), single constant force spring (Z, lift force: 1.5N)

  - System assembly (assembly, final test)

  - Factory Acceptance Test (functional test) shall include verification of travel range, resolution, material.

- Controller

  - The controllers shall be capable of controlling all stages. They shall also support remote operation via an Ethernet interface. The remote control interface shall include a PC-based control interface, as well as a C dynamic link library (*.DLL) and corresponding drivers.

- Inspection & Acceptance Criteria

  - The vendor shall provide factory acceptance test (FAT) documents for the positioners.

  - The visual inspection shall include verification of dimensions, connectors, and accessories.

  - Final inspection shall be carried out after functional testing at NSRRC. The functional tests shall include verification of travel range, moving speed, vacuum operation, and controller functionality.